## Slide 1: A Biaxial Tissue Stretcher
Client: Frank Yin, MD. Ph.D
Group 30
Joshua Leibowitz
Krista Vedvik
Christopher Zarins

## Slide 2: Need for a Biaxial Cell Stretcher
Studying the effects of mechanical force on aortic endothelial cells
Orientation and organization of cells depends on exact stretching qualities
Controlling deformation in both directions gives the most accurate and meaningful results

## Slide 3: Design Requirements
Parameter | Value
Maximum Strain | 40%
Strain Resolution | 0.50%
Maximum Strain Rate | 40% / s
Device Size | 19.5 in W x 19.5 in D x 23 in H
Operating Temperature | 37.5 ˚ C
Operating Humidity | 100%
Central Region of Uniform Strain | 1 cm x 1 cm
Cost | < $35,000

## Slide 4: Overview of Design Concept
Four linear actuators mounted above bath
Membrane rods for fixation mechanism

## Slide 5: Overview of Design Concept

## Slide 6: Drivetrain: Details

## Slide 7: Drivetrain: Assembly

## Slide 8: Drivetrain: Analysis
Force
  Required force:
  Actuator spec: 75 lb max
Resolution
  0.5% during 0.5 in. stretch=0.0025 in
  Actuator spec: 0.0004 in/step
Speed
  0.5 in/sec for 40% stretch/sec
  Actuator spec: 3.3 in/sec

## Slide 9: Drivetrain: Analysis
Aluminum 6061 Alloy
  For machined parts
  Corrosion resistant
  Relatively inexpensive
  Reduces total weight by 20% compared to steel

## Slide 10: Membrane Fixation: Details

## Slide 11: Membrane Fixation: Assembly

## Slide 12: Membrane Punch

## Slide 13: Membrane Fixation: Analysis
Membrane rods
  Middle ground between sutures and clamps
  Mitigate fluid shear
Finite Element Simulations

## Slide 14: Membrane Fixation: Analysis
Membrane Rods
  Some resolution decrease due to bending
Euler-Bernoulli Theorem

## Slide 15: Parts and Manufacturing
Drivetrain Manufacturing Specifics |  |  |  |  |  | 
Part | Manufacturer | Part Number | Quantity | Ind. Cost | Ind. Weight (lb) | Lead Time
End Plate | Firstcut | Quote # 253294 | 4 | $218 | 0.68764 | 5-14 bus. days
Linear Bearings | McMaster-Carr | 64825K11 | 8 | $59.51 | 0.4 | 3-7 bus. days
Linear Rails | McMaster-Carr | 1240K51 | 8 | $51.12 | 0.335 | 3-7 bus. days
Linear Actuator | Ultramotion | D-A.063-IPX17-2-/4-ESD | 4 | $875 | 1.1 | 8 weeks
Stepper Controller | Ultramotion | ST5-Q | 4 | $360 | 0 | 8 weeks
Stepper Controller Power Supply | McMaster-Carr | 7010K71 | 1 | $277.27 | 0 | 3-7 bus. days
Terminal Block | McMaster-Carr | 7527K45 | 1 | $2.35 | 0 | 3-7 bus. days
Main Rod | Firstcut | Quote # 253294 | 4 | $126 | 0.08852 | 5-14 bus. days
Total |  |  |  | $7,481 | 13.38464 | 
Membrane/Membrane Fixation Manufacturing Specifics |  |  |  |  |  | 
Part | Manufacturer | Part Number | Quantity | Ind. Cost | Ind. Weight (lb) | Lead Time
Standoff | Firstcut | Quote # 253294 | 4 | $88 | 0.06459 | 5-14 bus. days
Mounting Block | Firstcut | Quote # 253294 | 4 | $132 | 0.21941 | 5-14 bus. days
Membrane Rods | McMaster-Carr | 88915K37 | 1 | $10.88 | 0.335 | 3-7 bus. days
Bottom Rod Support | Firstcut | Quote # 253294 | 4 | $78 | 0.00902 | 5-14 bus. days
Needles | Vita Needle | N/A | 1 | $30 | 0 | 1-2 weeks
Alignment Rods | McMaster-Carr | 4416T12 | 1 | $6.27 | 0 | 3-7 bus. days
Top Punch | Firstcut | Quote # 253294 | 1 | $169 | 0 | 5-14 bus. days
Bottom Punch | Firstcut | Quote # 253294 | 1 | $189 | 0 | 5-14 bus. days
Total |  |  |  | $1,597 | 1.50708 | 

## Slide 16: Parts and Manufacturing
Fastener Manufacturing Specifics |  |  |  |  |  | 
Part | Manufacturer | Part Number | Quantity | Pkg. Cost | Ind. Weight (lb) | Lead Time
Frame to end plate | McMaster-Carr | 92185A589 | 16 | $5.75 | 0.047 | 3-7 bus. days
Linear bearing to end plate | McMaster-Carr | 96710A321 | 32 | $10.43 | 0.005 | 3-7 bus. days
End plate to L.A. | McMaster-Carr | 92703A211 | 16 | $8.21 | 0.001 | 3-7 bus. days
Into threaded end of rail | McMaster-Carr | 92703A408 | 8 | $6.96 | 0.01 | 3-7 bus. days
Into threaded end of L.A. | McMaster-Carr | 91771A557 | 4 | $9.60 | 0.009 | 3-7 bus. days
Mounting block to main rod | McMaster-Carr | 92185A548 | 8 | $6.67 | 0.029 | 3-7 bus. days
Membrance Spacers | McMaster-Carr | 94639A466 | 128 | $6.13 | 0.00012 | 3-7 bus. days
Nut (support bottom of membranes) | McMaster-Carr | 90730A001 | 8 | $6.55 | 0.00029 | 3-7 bus. days
Total |  |  |  | $108 | 1.29368 | 
Miscellaneous Parts Manufacturing Specifics |  |  |  |  |  | 
Part | Manufacturer | Part Number | Quantity | Pkg. Cost | Ind. Weight (lb) | Lead Time
Frame (entire structure) | 8020 | N/A | 1 | $750 | 16 | 3-7 bus. days
Laboratory Scissor Jack | Home Science Tools | CE-JACK8X8 | 1 | $44.95 | 0 | 7-10 bus. Days
Bath | Dynalon Labware | 451005 | 1 | $13.34 | 0 | 3-7 bus. days
Total |  |  |  | $808.29 | 16 | 
Total Weight: 32.19 lb
Total Projected Cost: $9,994.29

## Slide 17: Safety
Relatively low-risk device
  Only used in laboratory setting
  Trained lab assistants
Risk of pinching/poking fingers
  Membrane punch
  Assembly
Risk of electrical shock
  Controller power supply
  Degradation of linear actuator wires

## Slide 18: Conclusions
Objectives met
  Capable of 40% stretch at 40%/sec
  Capable of double the prescribed resolution
  Fits in incubator/humidity permitting
  Within budget
Parameter | Constraint | Actual Value
Maximum Strain | 40% | 100%
Strain Resolution | 0.50% | >0.26%
Maximum Strain Rate | 40% / s | 264% / s
Device Size | 19.5 in W x 19.5 in D x 23 in H | 18.6 in W x 18.6 in D x 22 in H
Operating Temperature | 37.5 ˚ C | 37.5 ˚ C
Operating Humidity | 100% | 100%
Central Region of Uniform Strain | 1 cm x 1 cm | 1 cm x 1 cm
Cost | < $35,000 | ~ $10,000

## Slide 19: Conclusions
No IP
  No mass production
Future Directions
  Prototyping and troubleshooting
  Precise calibration
  Better vertical access
  Complex FES
Design is a dynamic process!!

## Slide 20: References
Balland, M., et. al. Power Laws in Microrheology Experiments on Living Cells: Comparative Analysis and Modeling. Phys. Rev. E Stat. Nonlinear Soft Matter Phys. 74, (2006)
Collinsworth, A. et. al. Apparent Elastic Modulus and Hysteresis of Skeletal Muscle Cells Throughout Differentiation. Am. J. Cell Physiol. 283, C1219-C1227 (2002)
McGarry, J. et. al. A Comparison of Strain and Fluid Shear Stress in Stimulating Bone Cell Responses- A Computational and Experimental Study. FASEB J. 19, 482-484 (2005)
Thompson, M. et. al. Quantification and Significance of Fluid Shear Stress Field in Biaxial Cell Stretching Device. Biomech. Model Mechanobiol. 10, 559-564 (2011)
Yin, F., Chew, P., Zeger, S. An Approach to Quantification of Biaxial Tissue Stress-Strain Data. J. Biomech. 19, 27-37 (1986)
Zeng, D. et. al. Young’s Modulus of Elasticity of Schlemm’s Canal Endothelial Cells. Biomech. Model Mechanobiol. 9, 19-33 (2010)

## Slide 21: Questions?